SEWA AI PROVISIONAL PATENT APPLICATION — VERSION 2.0

Sovereign Multi-Engine Clinical AI Platform for Automated Diagnostic Triage and Safety Assurance

Application Type: UK Provisional Patent Application

Filing Body: UK Intellectual Property Office (UKIPO)

Applicant: Storegrill Inc Ltd

Registered Address: 18 Hampshire Close, Binley, Coventry, CV3 2FP, England

Company No.: 14581073

Inventor: Dolapo Ariyo

Priority Date: March 2026

Version: v2.0 · March 2026

Contact: info@sewaxai.health · www.sewaxai.health

Tagline: from hours, to seconds

ABSTRACT

A multi-engine clinical artificial intelligence platform and method for automated diagnostic triage and safety assurance in medical imaging workflows. The system ingests medical imaging studies (CT, MRI, X-ray, Ultrasound) via DICOM, FHIR R4, or HL7v2 and routes each study through four sequentially coupled AI engines: (1) a compound-scaled convolutional neural network triage engine detecting up to fourteen life-threatening conditions in under 620 milliseconds; (2) a stochastic Monte Carlo Dropout uncertainty engine decomposing prediction variance into epistemic and aleatoric components; (3) a vision-language model generating structured, editable four-section radiology reports in eight languages; and (4) a nine-layer metacognition engine performing automated self-audit and hallucination-risk scoring before clinical alert dispatch. The platform further comprises a federated learning subsystem enabling participating institutions to collaboratively improve the shared model using differentially private gradient updates without transmitting protected health information, and a sovereign deployment architecture operable in fully offline edge environments without internet connectivity. All clinical outputs require qualified human sign-off enforced as a database-level constraint, delivering door-to-alert latency under two seconds.

1. FIELD OF THE INVENTION

The present invention relates to artificial intelligence applied to clinical diagnostic imaging. Specifically, it concerns a multi-engine orchestrated software platform that automates radiology triage, quantifies AI prediction uncertainty, autonomously self-audits for hallucinations, generates structured diagnostic reports using vision-language models, delivers real-time multi-channel clinical alerts, and operates across sovereign cloud and offline edge deployments within a statutory human-in-the-loop framework.

2. BACKGROUND

2.1 Global Radiologist Deficit

The global healthcare system faces a structural and worsening radiologist workforce deficit. In the United Kingdom, the radiologist vacancy rate has reached 29% of established posts (Royal College of Radiologists, 2025), resulting in 1.5 million backlogged imaging reports annually. In sub-Saharan Africa, radiologist density is below 0.4 per million population against a WHO-recommended minimum of 20, representing an effective 98%+ deficit in Nigeria, Kenya, and Ghana. Patients die not from their conditions but from diagnostic delays.

2.2 Deficiencies of Prior Art

Existing clinical AI systems present six material deficiencies:

(a) Single-condition scope. Prior art systems detect one condition per deployment (e.g., large-vessel occlusion stroke, or intracranial haemorrhage only). Emergency departments require simultaneous multi-condition detection within a single inference pipeline.

(b) Absence of uncertainty quantification. All prior art systems output a single point-estimate confidence score without quantifying prediction uncertainty. A model expressing high numerical confidence in an incorrect prediction provides no mechanism for a clinician to distinguish a reliable from an unreliable prediction.

(c) Absence of automated self-audit. No prior art system audits its outputs before clinical delivery. AI hallucinations — plausible but factually incorrect outputs — are a known failure mode of all neural network architectures. No prior art provides a technical gate preventing hallucinated critical findings from reaching clinicians.

(d) Black-box outputs. Existing systems do not generate structured human-readable diagnostic reports, nor do they provide spatial attribution of AI decision-driving features.

(e) Cloud dependency. All prior art clinical AI systems require persistent high-bandwidth internet connectivity. This renders them inoperable in the 70%+ of low-and-middle-income clinical settings lacking reliable internet infrastructure.

(f) No human-in-the-loop guarantee. Prior art systems permit autonomous AI outputs to be acted upon without verified human review, inconsistent with EU AI Act Article 14 and NHS DCB0129.

2.3 Technical Need

There exists a clear and unmet need for a clinical AI system that: (i) detects multiple acute conditions from a single submission within clinically actionable latency; (ii) quantifies its own prediction uncertainty; (iii) audits its outputs for factual consistency before alert dispatch; (iv) generates structured, editable diagnostic reports with spatial attribution; (v) operates in sovereignty-preserving, disconnected environments; (vi) implements federated learning without transmitting protected health information; and (vii) enforces a statutory human-in-the-loop guarantee on all clinical outputs.

3. SUMMARY OF THE INVENTION

The invention provides a computer-implemented multi-engine clinical AI platform comprising nine interconnected subsystems:

[S1] Scan ingestion and pre-processing: Receives medical imaging data via DIMSE C-STORE, FHIR R4, HL7v2 MLLP, or direct upload; applies modality-specific Hounsfield Unit windowing and pixel normalisation; produces a standardised imaging tensor for GPU inference.

[S2] Triage engine: EfficientNet-B4 convolutional neural network detecting up to fourteen clinical conditions with urgency classification (CRITICAL/URGENT/NORMAL/ROUTINE) in median 620ms, producing Grad-CAM heatmaps and SHAP feature attribution simultaneously.

[S3] Uncertainty engine: Monte Carlo Dropout over twenty stochastic forward passes, decomposing prediction variance into epistemic and aleatoric uncertainty; mandatory human escalation when composite uncertainty exceeds a configurable threshold, independent of triage confidence.

[S4] VLM report generator: Vision-language model producing structured four-section radiology reports (Technique / Findings / Impression / Recommendation) in eight languages, with clinician edit tracking for continuous learning.

[S5] Metacognition engine: Nine sequential validation layers evaluating modality-finding plausibility, urgency-finding consistency, confidence floors, SHAP coherence, uncertainty-confidence conflict, report-finding concordance, demographic plausibility, composite hallucination risk scoring, and producing a binary safe-to-alert gate.

[S6] Multi-channel alert subsystem: Simultaneous FCM push, WebSocket, email, and HL7v2 ORU^R01 dispatch upon metacognition PASS; Stroke Code pathway triggering additional HL7v2 ADT^A02 and ORM^O01 messages.

[S7] Human-in-the-loop enforcement: Database-level constraint requiring qualified clinician digital signature before clinical output delivery; enforced independently of application logic.

[S8] Federated learning subsystem: Differentially private gradient aggregation enabling collaborative model improvement without PHI transmission; gated by human approval before production deployment.

[S9] Sovereign deployment architecture: Cloud-native AKS GPU inference path and offline INT8-quantised edge path with asynchronous synchronisation; Merkle-chained tamper-evident audit trail across both paths.

4. DETAILED DESCRIPTION

4.1 System Architecture — Figure 1

The SEWA AI platform implements a four-engine sequential pipeline for clinical diagnostic triage and safety assurance. Medical imaging studies are received via DICOM C-STORE, FHIR R4, HL7v2 MLLP, or direct upload API into the scan ingestion subsystem [S1]. Pre-processing applies modality-specific Hounsfield Unit windowing and pixel normalisation to produce a standardised multi-channel imaging tensor [1,3,512,512] for GPU inference.

Engine 1 (Triage) comprises an EfficientNet-B4 compound-scaled convolutional neural network detecting up to fourteen clinical conditions across CT Head, CT Chest, X-Ray, MRI Brain, and multi-modality inputs. Inference executes in 380–620ms, producing per-condition probability scores, urgency classification (CRITICAL/URGENT/NORMAL/ROUTINE), Grad-CAM spatial heatmaps, and SHAP feature attribution.

Engine 2 (Uncertainty) executes N=20 stochastic forward passes through the triage engine with Dropout layers active (Monte Carlo Dropout). Prediction variance is decomposed into epistemic uncertainty (model uncertainty) and aleatoric uncertainty (data noise). A composite uncertainty score U = 0.6*sigma²_ep + 0.4*sigma²_al is computed. If U exceeds a configurable threshold (default 0.25), the case is mandatorily routed to human clinical review independently of downstream engines.

Engine 3 (VLM Report Generator) accepts triage outputs, SHAP top-5 spatial narrative hints, patient metadata, and DICOM parameters as inputs. A CLIP medical vision encoder produces a [1,512] image embedding, which is decoded by an LLM with LoRA clinical fine-tuning into a structured four-section radiology report: Technique, Findings, Impression, and Recommendation. Reports are generated in eight languages and are clinician-editable with diff tracking.

Engine 4 (Metacognition) implements nine sequential validation layers evaluating AI outputs before CRITICAL alert dispatch. The layers assess: (1) modality-finding plausibility; (2) urgency-finding consistency with auto-correction; (3) confidence floors per urgency class; (4) SHAP coherence between attribution and finding; (5) uncertainty-confidence conflict detection; (6) report-finding concordance via negation detection; (7) demographic plausibility against epidemiological priors; (8) composite hallucination risk score H as a weighted combination of signals from Layers 1–7; and (9) binary safe-to-alert gate permitting dispatch only when no REJECT from Layers 1–7 and H <= 0.30.

Upon metacognition PASS, the multi-channel alert subsystem [S6] dispatches simultaneously via FCM push notification, WebSocket event, email, and HL7v2 ORU^R01. For Stroke Code conditions (ICH, LVO, SAH, Aortic Dissection), additional HL7v2 ADT^A02 and ORM^O01 messages are triggered. All clinical outputs require qualified clinician digital signature before delivery, enforced as a database-level constraint [S7].

4.2 Triage Engine

[4.2.1] The triage engine comprises EfficientNet-B4 (18.9M parameters), pre-trained on ImageNet-21k and fine-tuned on >=500,000 labelled medical imaging studies using transfer learning with differential learning rates and class-weighted binary cross-entropy loss.

[4.2.2] Pre-processing applies three simultaneous windowing presets producing a [1,3,512,512] three-channel tensor. For CT head: brain (W80/L40), subdural (W130/L75), bone (W2800/L600). For CT chest: lung (W1500/L-600), mediastinal (W350/L25), bone.

[4.2.3] Fourteen conditions are detected across modalities, with CRITICAL thresholds as follows:

| Modality | Conditions | CRITICAL Threshold |
| --- | --- | --- |
| CT Head | ICH, Ischaemic Stroke, SDH, SAH | >= 0.45 |
| CT Chest | PE, Pneumothorax, Aortic Dissection | >= 0.45 |
| X-Ray | Pneumonia, TB, Pneumothorax, Cardiomegaly | >= 0.50 |
| MRI Brain | Ischaemic Stroke (DWI), Space-Occupying Lesion | >= 0.50 |
| Multi | Bone Fracture | >= 0.50 |

[4.2.4] Grad-CAM heatmaps are computed simultaneously with inference by backpropagating the gradient of the predicted class score through the final convolutional feature maps, applying ReLU, and upsampling to input dimensions.

[4.2.5] SHAP values are computed via Kernel SHAP on superpixel segmentation; top-five contributing regions are passed as structured narrative hints to the VLM.

4.3 Uncertainty Engine

[4.3.1] The uncertainty engine executes N=20 stochastic forward passes through the triage engine with Dropout layers active (MC-Dropout). For condition c with predictions P={p1,...,pN}:

Epistemic uncertainty: sigma²_ep = (1/N) * Sum(i=1 to N) of (pi_c - mu_c)²

Aleatoric uncertainty: sigma²_al = (1/N) * Sum(i=1 to N) of pi_c * (1 - pi_c)

Composite score: U = 0.6 * sigma²_ep + 0.4 * sigma²_al

[4.3.2] If U > 0.25 (default threshold, configurable), the case is mandatorily routed to human clinical review independently of the metacognition engine. This rule cannot be overridden by downstream engines.

[4.3.3] Out-of-distribution detection is computed as the Mahalanobis distance between the input embedding and the training distribution mean, flagging OOD inputs.

4.4 VLM Report Generator — Figure 2

The VLM report generator accepts four input categories: (1) triage output (findings, confidence, urgency); (2) SHAP top-5 feature descriptions as spatial narrative hints; (3) patient metadata (age, sex, indication); and (4) DICOM parameters (modality, protocol, contrast).

A CLIP Medical Vision Encoder processes the imaging tensor to produce a [1,512] image embedding. This embedding is fed into an LLM Decoder with LoRA clinical fine-tuning, which performs structured generation of a four-section radiology report:

TECHNIQUE: Acquisition parameters and scan protocol.
FINDINGS: Region-by-region description with SHAP-attributed spatial references.
IMPRESSION: Primary conclusion combined with urgency classification.
RECOMMENDATION: Clinical action and follow-up instructions.

The generated report undergoes Metacognition Layer 6 concordance checking. The clinician-facing interface presents the report in an editable format with diff tracking. Upon radiologist digital signature, the report is converted to FHIR DiagnosticReport format and dispatched via HL7 ORU^R01.

4.5 Metacognition Engine — Nine-Layer Architecture — Figure 3

The metacognition engine is the core safety innovation. Nine sequential layers evaluate AI outputs before CRITICAL alert dispatch. No AI result is ever silently dropped; any FAIL routes to human review with full AI findings displayed.

Layer 1: Modality-Finding Plausibility. Evaluates whether a detected condition is anatomically plausible for the imaging modality. Example: TB on CT Head triggers immediate REJECT.

Layer 2: Urgency-Finding Consistency. Evaluates whether the urgency classification matches the detected finding. Example: ICH marked NORMAL triggers AUTO-CORRECT to CRITICAL plus FLAG.

Layer 3: Confidence Floor by Urgency. Enforces minimum confidence thresholds per urgency class. CRITICAL findings with confidence < 0.45 are FLAGged and routed to human queue.

Layer 4: SHAP Coherence Check. Evaluates whether the top attribution region is anatomically consistent with the detected finding. Coherence score < 0.5 triggers FLAG.

Layer 5: Uncertainty-Confidence Conflict. Detects contradictory signals where confidence > 0.85 AND uncertainty > 0.25, indicating potential model overconfidence on unreliable data.

Layer 6: Report-Finding Concordance. Checks concordance between the natural language impression and the urgency classification via negation detection. Example: VLM states 'no acute finding' but urgency = CRITICAL triggers FLAG.

Layer 7: Demographic Plausibility. Evaluates findings against epidemiological priors for patient demographics. Example: PE in patient < 10 years triggers FLAG with annotation.

Layer 8: Composite Hallucination Risk Score H. Computed as a weighted combination of signals from Layers 1–7: H = w1*(1-coherence) + w2*conflict + w3*uncertainty + w4*OOD_distance + w5*negation_detected. If H > 0.30, autonomous alert is BLOCKed and the case is routed to human queue.

Layer 9: Safe-to-Alert Gate. Produces binary PASS or FAIL decision. PASS permits multi-channel alert dispatch. FAIL always routes to human review — the AI result is never discarded.

Key property: No AI result is ever silently dropped. FAIL routes to human review with full AI findings displayed, ensuring no critical finding is lost.

4.6 Stroke Code Activation — Figure 4

Upon metacognition PASS for a time-critical condition (ICH, LVO, SAH, or Aortic Dissection) matching the Stroke Code condition set, six channels are dispatched in parallel within < 5 seconds from PASS:

[1] FCM Push to all radiologists and on-call doctor.
[2] WebSocket event to all active browser sessions.
[3] Email to radiologist, referring doctor, and on-call clinician.
[4] HL7 ORU^R01 to EMR with STAT priority.
[5] HL7 ADT^A02 stroke unit transfer order.
[6] HL7 ORM^O01 CTA head+neck order (STAT).

The door-to-alert timestamp is recorded in milliseconds in an immutable audit log. All events are SHA-256 Merkle-chained.

The reviewing clinician is presented with DICOM imaging, Grad-CAM heatmap, SHAP attribution, and AI summary. Upon review, the radiologist either confirms the AI findings (signs report, generates FHIR DiagnosticReport) or overrides (documents reason, signs own report). Both outcomes generate training signals captured for continuous learning.

4.7 Federated Learning Subsystem — Figure 5

The federated learning subsystem enables privacy-preserving collaborative model improvement across participating institutions without transmitting protected health information.

Protocol:
(a) The central aggregation server distributes global model weights W_global to all participating institutions.
(b) Each institution performs local training on locally-held data; no PHI leaves the institution.
(c) Local model updates are computed as delta_Wi = W_local - W_global.
(d) Differential privacy noise is applied to local updates: (epsilon=1.0, delta=10^-5) before transmission.
(e) Only differentially private gradient updates delta_Wi' are transmitted to the aggregation server.
(f) Weighted averaging computes the proposed global model: W_global_new = Sum(i) of (ni/N) * delta_Wi'.
(g) A human approval gate requires digital sign-off from a designated Chief Scientific Officer (CSO) before production deployment.
(h) All approval events are audit-logged in a tamper-evident chain.

Anomaly detection is applied to received gradient updates to detect potential poisoning attacks or corrupted local models.

4.8 Audit Trail — Figure 6

All clinical events are recorded in a tamper-evident audit log with the following structure per record:

event_id: UUID v4
timestamp: ISO 8601 UTC (microsecond precision)
user_id: UUID of authenticated clinician
event_type: SCAN_RECEIVED | AI_RESULT | ALERT_SENT | REPORT_SIGNED | OVERRIDE | FEDERATED_UPDATE
payload_hash: SHA-256 hash of resource at event time
prev_hash: SHA-256 hash of previous audit record (chain linkage)
merkle_root: batch Merkle root (nightly recompute)

Hash chain integrity ensures tampering with any record breaks the chain. Integrity is verified nightly, with alerts generated on mismatch. Retention period is 7 years per NHS Records Management Code 2021.

4.9 Sovereign Deployment Architecture

Cloud Path (Azure Kubernetes Service):

CPU node pool: API gateway, orchestration, database proxy, monitoring (Standard_D4s_v3, 4vCPU/16GB, minimum 3 nodes).

GPU node pool: All four AI engines as containerised microservices (Standard_NC4as_T4_v3, T4 16GB, minimum 2 nodes).

PostgreSQL 15 with row-level tenant isolation, Redis 7, Azure Blob with AES-256-GCM encryption.

99.9% uptime SLA with blue-green zero-downtime deployments.

Offline Edge Path:

INT8-quantised triage model (~18MB, ~80ms inference on CPU).

Offline SQLite database with local cache capacity for 500 scans.

SMS fallback via GSM modem (Africa's Talk / Twilio SMS API).

Asynchronous cloud synchronisation when connectivity is restored.

Merkle-chained audit preserved across offline and online periods.

The device provides full AI triage and alert capability in both connected and disconnected operation modes.

5. CLAIMS

Claim 1 — Multi-Engine Diagnostic Orchestration (Method)

A computer-implemented method for clinical diagnostic triage and safety assurance in medical imaging, comprising:

(a) receiving, by a scan ingestion subsystem, a medical imaging study via any of DICOM C-STORE, FHIR R4, HL7v2 MLLP, or direct upload API;

(b) pre-processing the imaging data by applying modality-specific Hounsfield Unit windowing and producing a standardised multi-channel imaging tensor;

(c) executing, by a triage engine comprising a compound-scaled convolutional neural network, inference producing per-condition probability scores for at least two predefined clinical conditions and an urgency classification;

(d) executing, by an uncertainty engine, a plurality of stochastic forward passes through the triage engine with dropout layers active, computing a quantified uncertainty score comprising at least an epistemic uncertainty component;

(e) generating, by a vision-language model, a structured four-section diagnostic report from triage outputs and patient metadata;

(f) evaluating, by a metacognition engine comprising at least nine sequential validation layers, triage and report outputs against clinical plausibility rules, uncertainty-confidence consistency, and report-finding concordance, producing a binary safe-to-alert decision;

(g) dispatching, upon a PASS safe-to-alert decision, clinical alerts simultaneously via at least mobile push notification and HL7v2 ORU^R01; and

(h) requiring a qualified clinician's digital signature before clinical output delivery, enforced as a database-level constraint independent of application logic.

Claim 2 — Metacognition Safety Layer (Method)

A method for automated safety assurance of AI-generated clinical diagnostic outputs through nine sequential validation layers, wherein:

Layer 1 evaluates modality-finding plausibility;

Layer 2 evaluates urgency-finding consistency with auto-correction;

Layer 3 enforces minimum confidence floors per urgency class;

Layer 4 computes SHAP coherence between attribution and finding;

Layer 5 detects conflicts between high confidence and high uncertainty;

Layer 6 checks concordance between natural language impression and urgency via negation detection;

Layer 7 evaluates demographic plausibility against epidemiological priors;

Layer 8 computes composite hallucination risk score as weighted combination of signals from Layers 1-7;

Layer 9 produces binary safe-to-alert decision permitting dispatch only when no REJECT from Layers 1-7 and hallucination risk <= 0.30;

wherein any failure routes the case to human clinical review rather than discarding the AI finding.

Claim 3 — Federated Learning with Human Approval Gate (System)

A system for privacy-preserving collaborative clinical AI model training, comprising:

(a) a central aggregation server distributing global model weights to participating institutions;

(b) local training at each institution on locally-held data without PHI leaving the institution;

(c) application of differential privacy noise to local model updates before transmission;

(d) transmission of only differentially private gradient updates to the aggregation server;

(e) weighted averaging of received updates to compute a proposed global model update;

(f) a human approval gate requiring digital sign-off from a designated approver before production deployment; and

(g) audit logging of all approval events in a tamper-evident chain.

Claim 4 — Stroke Code Activation Workflow (Method)

A method for automated stroke code activation comprising, upon metacognition PASS for a time-critical condition:

simultaneously dispatching via (i) mobile push notification, (ii) WebSocket event, (iii) HL7v2 ADT^A02 patient transfer message, (iv) HL7v2 ORM^O01 imaging order, and (v) HL7v2 ORU^R01 result message; recording door-to-alert time in milliseconds in an immutable audit log; and presenting AI findings with Grad-CAM, SHAP, and uncertainty score to the reviewing clinician.

Claim 5 — Sovereign Self-Hosted Deployment (Device)

A device for sovereign self-hosted clinical AI deployment comprising:

(a) a cloud-native path with GPU-enabled container orchestration, tenant-isolated encrypted database, and web application firewall; and

(b) an offline edge path with INT8-quantised AI model operable without internet connectivity, local audit log with Merkle chain integrity maintained across offline periods, and SMS alert fallback via GSM modem;

wherein the device provides full AI triage and alert capability in both connected and disconnected operation modes.

Claim 6 — Uncertainty-Gated Human Escalation (Method)

A method wherein a composite uncertainty score derived from epistemic and aleatoric decomposition of Monte Carlo Dropout variance mandatorily routes a medical imaging case to human clinical review when the score exceeds a configurable threshold, independently of the triage engine confidence score and independently of the metacognition engine, such that human escalation cannot be bypassed by any downstream AI component.

Claim 7 — Dependent: Three-Channel CT Pre-Processing

The method of Claim 1, wherein pre-processing for CT inputs applies three simultaneous window presets producing a three-channel tensor including at least a brain window and a subdural window for head CT, and a lung window and mediastinal window for chest CT.

Claim 8 — Dependent: Immutable Audit Trail

The system of Claim 3 or Claim 5, wherein all clinical events are recorded in a tamper-evident audit log using SHA-256 hash chaining wherein each record contains the hash of the previous record, with Merkle tree batch verification computed at regular intervals, maintained across both cloud and offline edge deployments.

Claim 9 — Dependent: Multi-Language VLM

The method of Claim 1, wherein the vision-language model generates structured diagnostic reports in a plurality of natural languages comprising at least English, French, Arabic, and Swahili, with language selection determined by clinical tenant configuration.

Claim 10 — Dependent: Hallucination Risk Score

The method of Claim 2, wherein Layer 8 composite hallucination risk score computation combines: a SHAP coherence deficit from Layer 4; an uncertainty-confidence conflict indicator from Layer 5; the composite uncertainty score from the uncertainty engine; a Mahalanobis out-of-distribution distance; and a negation detection indicator from Layer 6 analysis of the VLM impression section.

6. ANNEXES

Annex A — Technical Infrastructure

| Component | Specification | Purpose |
| --- | --- | --- |
| AKS CPU Pool | Standard_D4s_v3 (4vCPU/16GB), x3 min | API, orchestration, DB proxy |
| AKS GPU Pool | Standard_NC4as_T4_v3 (T4 16GB), x2 min | AI inference (4 engines) |
| PostgreSQL | Azure Flexible Server v15, Zone-HA | Clinical data, row-level tenant security |
| Redis | Azure Cache for Redis v7 | JWT blacklist, rate limiting, AI cache |
| Blob Storage | Azure Blob, AES-256-GCM | DICOM, Grad-CAM, reports |
| Edge Node Min. | 4-core CPU, 8GB RAM, 128GB SSD | INT8 offline inference |
| Edge Model | EfficientNet-B4 INT8 quantised, ~18MB | Offline triage |
| Edge Alert | GSM modem, Africa's Talk SMS API | Offline critical alerts |

Annex B — Sample Workflow Data (Anonymised)

{
  "doorToAlertMs": 1842,
  "triageResult": {
    "finding": "ICH",
    "confidence": 0.942,
    "urgency": "CRITICAL",
    "inferenceMs": 384
  },
  "uncertaintyResult": {
    "epistemic": 0.031,
    "aleatoric": 0.058,
    "composite": 0.042,
    "highFlag": false
  },
  "metacognitionResult": {
    "layer1": "PASS",
    "layer2": "PASS",
    "layer3": "PASS",
    "layer4": "PASS",
    "layer5": "PASS",
    "layer6": "PASS",
    "layer7": "PASS",
    "layer8_H": 0.08,
    "layer9": "PASS"
  },
  "alertChannels": {
    "fcm": "delivered_312ms",
    "websocket": "delivered_8ms",
    "hl7_oru": "delivered_445ms",
    "stroke_code": true
  }
}

Annex C — Clinical Safety Risk Summary

| Risk | Pre-Control | Controls Applied | Post-Control |
| --- | --- | --- | --- |
| False negative (missed ICH) | High (RPN 10) | Uncertainty escalation; metacognition; human sign-off | ALARP (RPN 5) |
| AI hallucination | High (RPN 8) | Metacognition 9-layer; H <= 0.30 gate | Negligible (RPN 4) |
| Automation bias | Medium (RPN 12) | DB-level HITL constraint; training | ALARP (RPN 8) |
| PHI via federated learning | High (RPN 8) | Differential privacy; no PHI leaves institution | ALARP (RPN 4) |
| Offline alert failure | High (RPN 12) | SMS fallback; offline audit; async sync | ALARP (RPN 6) |

7. STATEMENT OF INVENTORSHIP

I, Dolapo Ariyo, as sole inventor, conceived and implemented the four-engine AI pipeline architecture, federated learning subsystem with differential privacy, sovereign deployment architecture (cloud and edge), Merkle-chained audit trail, multi-channel alert dispatch system, Stroke Code activation workflow, human-in-the-loop enforcement subsystem, and the complete technical framework of the SEWA AI platform described in this application.

Signed: _______________________________

Dolapo Ariyo, Inventor

Date: March 2026

Address: c/o Storegrill Inc Ltd, 18 Hampshire Close, Binley, Coventry, CV3 2FP